## Sheet1
Failure mode | Area | Likelihood | Impact | Mitigation | Comments | Acceptance
Malfunction causes the pump to deliver boluses in an uncontrolled manner. | Medtronic pump | Low | High | Device validation performed by Medtronic. | This risk is inherent in the current pump design, and has been mitigated by Medtronic's validation process. This risk has been accepted by the FDA. | Not applicable. Risk already accepted by Medtronic and the FDA.
Pump user executes too many boluses due to confusion or incorrect training. | Human factor | Medium | High | User training via diabetes care professionals worldwide. Bolus Wizard to assist in making bolus decisions. | This risk is inherent in the current pump design and will be further mitigated by the use of an APS-based control algorithm which can be monitored and improved over time. | Not applicable. Risk exists regardless of whether an APS is additionally used for bolus.
Sensor readings drift from actual BG over time. Sensor readings are not reliable for making bolus decisions. | CGM | High | High | Calibrations 4 times per day; no bolus delivered if calibration is older than 7 hours; no bolus delivered if sensor reading changes significantly (2 mmol/L) from one reading to the next. | Based on user feedback as well as some studies, this risk is higher with a Medtronic CGM than with a Dexcom CGM.
Short circuit or other condition on the remote control causes continuous button commands to be sent - (single button) | APS hardware - remote control | Medium | None | It is not possible to activate the pump by pressing a single button.
Short circuit or other condition on the remote control causes continuous button commands to be sent - (multiple buttons) | APS hardware - remote control | Low | High | Each button is managed via an independent RPi GPIO pin and an independent optocoupler. The likelihood of failure across multiple buttons simultaneously is lower than the case of a single button failure. In the event that such an error does occur and multiple buttons send erroneous commands, then the following mitigation remains in place: Correct button sequence needed User notification on bolus Easy bolus setup 0.1 units per button press User may switch off the APS and continue managing diabetes manually via CGM and pump.
Lack of communication; current glucose and/or IOB unknown | APS hardware - Carelink USB | High | High | Overall design philosophy takes this high risk into account. Within the overall system, the primary devices for diabetes management remain the FDA-approved pump and the CGM, with user interaction. The APS is designed only to make decisions and send commands to the pump when it has sufficient recent information available, which it assesses to likely be reliable. Specifically, via software: In the absence of current IOB or glucose information from the pump, no action is taken. If communication breakdown is longer than 30 minutes, user is notified. If glucose moves outside user-set limits at any time, CGM alarms notify the user (CGM functionality).
Bolus sent does not take into account unmeasured factors e.g. exercise | APS software bolus algorithm | High | High | User may activate "exercise mode" via a pushbutton on the APS. Exercise mode reduces bolus to a set percentage of normal (e.g. 30%) which can be configured in software by the user. Carb recommendations & pump suspension as soon as glucose drop + IOB leads to a potential to drop below the target glucose level Low glucose alarms set on the pump To prevent overshoot of the target glucose level, no correction bolus is delivered when the glucose measures less than target + 1.5 mmol/L
Cumulative error or other condition causes repetitive bolus commands | APS software bolus algorithm | Medium | High | Controls implemented in software: No bolus can be sent without knowing current IOB. No bolus can lead to IOB above 3 units. If the algorithm recommends higher than 3u IOB, then the user is prompted to manually execute a bolus on the pump. No bolus sent if glucose is less than 7 mmol/L.
Bolus commands are not received by the pump; glucose levels rise |  | Low | Medium | Medtronic remote control demonstrates high reliability in transmitting commands. High glucose alarm is activated on the pump, notifying the user that a problem exists requiring manual attention.
Pump is suspended by the APS. Due to communication failure, the APS loses the capability to restart the pump |  | Medium | Medium | High glucose alarm is activated on the pump, notifying the user that a problem exists requiring manual attention.
IOB is measured moments prior to a manual bolus by user; system calculates and delivers a bolus nearly simultaneously to the user. |  | Medium | High | Bolus function checks last bolus time just prior to bolus. If no answer received, no bolus delivered. If last bolus is less than 4 minutes ago, no bolus delivered. Bolus function checks whether pump is currently bolusing. If yes, no bolus delivered. When user chooses to deliver a bolus (e.g. for a high carb meal), he/she only delivers the bolus via the bolus wizard on the pump, to ensure latest active insulin is always taken into account in the calculation.
Mitigations above fail; too much insulin is sent via bolus; glucose levels drop | APS software bolus algorithm | Low | High | Low glucose alarm is activated on pump at a relatively high glucose level, e.g. 5 mmol/L. Pump alarms alert user to the issue in order to take corrective action e.g. dextrose or glucagon.